WatchWalk Wearable mobility data collection

AX3 Set-up Guide

Instructions for configuring, deploying and retrieving Axivity AX3 sensors for WatchWalk processing.

| AUDIENCE Research staff and wearable technology teams | DEVICE Axivity AX3 | DOCUMENT Version 1.0 |
| --- | --- | --- |

| WORKFLOW Install, connect, record, deploy, retrieve, export | RECORDING WINDOW Seven-day continuous monitoring by default | OUTPUT FILE .CWA file renamed by participant and trial number |
| --- | --- | --- |

Before you start

Prepare the sensor, a Windows computer, the OMGUI/AX3 GUI software (free software), a micro-USB to USB-A data transfer cable and the participant or trial ID. Confirm local ethics, consent and data storage procedures before issuing the device.

Set-up workflow

| 01 Download AX3 GUI Download the AX3 GUI installer from the project repository using the green Download button. |  |
| --- | --- |

| 02 Install OMGUI Run the downloaded AX3-GUI-__.exe file and follow the installation prompts. The installer will install the Open Movement Graphical User Interface. |  |
| --- | --- |

| 03 Open OMGUI After installation, open OMGUI. Check that the main window opens correctly before connecting the sensor. |  |
| --- | --- |

| 04 Connect the sensor Plug the Axivity AX3 sensor into the computer using a micro-USB to USB-A data transfer cable. |  |
| --- | --- |

| 05 Confirm device detection When the sensor is linked to the computer, it should appear in the OMGUI device list. |  |
| --- | --- |

| 06 Open recording settings Select the sensor in the device list, then click Record. |  |
| --- | --- |

| 07 Review the recording window A Recording Settings pop-up window will open. Review the sensor ID and sampling settings before changing the schedule. |  |
| --- | --- |

| 08 Set interval recording Select Interval, then enter the Start Date, Start Time and Duration. Use the study-approved seven-day schedule unless the protocol states otherwise. |  |
| --- | --- |

| 09 Save and verify settings Click OK and confirm that the Recording session displayed in OMGUI is correct. |  |
| --- | --- |

Deployment, retrieval and export

| 10 Issue the device Give the configured sensor to the wearer together with the user instructions. Remind them to wear the device continuously for seven days unless discomfort or skin irritation occurs. Record the issue date, planned return date and device serial number in the study log. | Quality check |
| --- | --- |

| 11 Retrieve the sensor Collect the sensor after seven days and reconnect it to the computer using the data transfer cable. Check for signs of strap damage or device tampering before data download. | Quality check |
| --- | --- |

| 12 Copy and rename the .CWA file In File Manager, drag the .CWA file from the sensor to the study computer. Rename the file using the participant or trial number before secure transfer to the analysis team. A common convention is to use the first three digits for the participant number and the fourth digit for the trial number. For example, 0011.CWA. Participant codes with digits only (no letters) typically work best with the WatchWalk Platform |  |
| --- | --- |

| 13 Store and send securely Store the renamed .CWA file in the approved study folder, then send or upload it through the agreed secure pathway. Do not email identifiable participant information unless this is specifically approved in the study protocol. |  |
| --- | --- |

Quality checklist

| Stage | Check | Complete |
| --- | --- | --- |
| Installation | OMGUI opens without error on the study computer | ☐ |
| Connection | Sensor appears in the OMGUI device list | ☐ |
| Configuration | Start date, start time and duration match the study schedule | ☐ |
| Issue | Wearer receives the user instruction sheet and contact details | ☐ |
| Retrieval | Sensor is collected and reconnected after seven days | ☐ |
| Export | .CWA file is copied, renamed and stored in the approved study folder | ☐ |

Troubleshooting notes

| Issue | Recommended action |
| --- | --- |
| Sensor does not appear | Use a confirmed data transfer cable, try another USB port and reopen OMGUI. |
| Recording session looks incorrect | Cancel the settings, reopen Record and re-enter the interval details. |
| File name is unclear | Rename using the participant/trial convention before moving the file into the analysis folder. |